## Summary
Facility | DMI Hollywood | Manufacture | Siemens
Address 1 | 6517 Taft St | Model | Sensation 16
Address 2
City, State, ZIP | Hollywood, FL 33024 | Gantry SN | 50625
Room Number | CT | Clinical Engineering ID | NA
Accreditation Information
ACR CTAP | 52504-01 | Expiration | 2019-08-28
ACR Patient Type | Adult Only
Surveyor | Troy Zhou, PHD, DABR, DABSNM | Survey Date | 2020-06-26
Medical Physicist | Troy Zhou, PHD, DABR, DABSNM | Report Date | 2020-06-26
Medical Physicist Tests |  |  |  | Pass/Fail/NA
Review of CT Protocols |  |  |  | PASS
Scout Prescription Accuracy |  |  |  | PASS
Alignment Light Accuracy |  |  |  | PASS
Table Travel Accuracy |  |  |  | PASS
Radiation Beam Width |  |  |  | PASS
Low-Contrast Performance |  |  |  | PASS
Spatial Resolution |  |  |  | PASS
CT Number Accuracy |  |  |  | PASS
CT Number Uniformity |  |  |  | PASS
Artifact Evaluation |  |  |  | PASS
Dosimetry |  |  |  | PASS
Acquisition Display Calibration |  |  |  | PASS
Technologist QC Evaluation |  |  |  | Pass/Fail/NA
Water CT Number and SD (Daily) |  |  |  | PASS
Artifact Evaluation (Daily) |  |  |  | PASS
Wet Laser QC (Weekly) |  |  |  | PASS
Visual Checklist (Monthly) |  |  |  | PASS
Dry Laser QC (Monthly) |  |  |  | NA
Acquisition Display QC (Monthly) |  |  |  | PASS
Medical Physicist's Recommendations for Quality Improvement
None.

## Data
1. CT Protocol Review | [11 empty] | PASS
ACR Protocol |  | Adult Head |  | Adult Abd |  | CT Chest WO |  | Chest w Embolus |  | C-Spine WO |  | Sinus WO
RPID Code |  | RPID22 |  | RPID3 |  | RPID16 |  | RPID147 |  | RPID21 |  | RPID62
Facility Protocol |  | Routine Head |  | Abd/Pel WO |  | Routine Chest |  | PE Study |  | C/Spine WO |  | Sinus
kVp |  | 120 |  | 120 |  | 120 |  | 120 |  | 140 |  | 120
mA |  | 350 |  | 500 |  | 500 |  | 225 |  | 200 |  | 73
Rotation Time |  | 1 |  | 0.5 |  | 0.5 |  | 0.5 |  | 0.75 |  | 1
mAs |  | 350 |  | 250 |  | 250 |  | 112.5 |  | 150 |  | 73
Effective mAs |  | 350 |  | 250 |  | 200 |  | 150 |  | 300 |  | 73
Detector Width |  | 1.5 |  | 1.5 |  | 0.75 |  | 0.75 |  | 1.5 |  | 1.5
# Detectors |  | 12 |  | 16 |  | 16 |  | 16 |  | 16 |  | 16
Axial or Helical |  | Axial |  | Helical |  | Helical |  | Helical |  | Helical |  | Helical
Pitch |  | 1 |  | 1 |  | 1.25 |  | 0.75 |  | 0.5 |  | 1
Table Increment |  | 18 |  | 24 |  | 15 |  | 9 |  | 12 |  | 24
Scan FOV |  | 30 |  | 50 |  | 50 |  | 50 |  | 50 |  | 30
Dose Reduction |  | None |  | CAREDose 4D |  | CAREDose 4D |  | CAREDose 4D |  | None |  | None
Settings |  | NA |  | QRM200mAs |  | QRM200mAs |  | QRM150mAs |  | NA |  | NA
Recon Kernel |  | H31s |  | B30f |  | B30f |  | B41f medium+ |  | B70s Very Sharp |  | H70h Very Sharp
Recon Width |  | 4.5 |  | 5 |  | 2 |  | 3 |  | 2 |  | 2
Recon Interval |  | 4.5 |  | 5 |  | 2 |  | 3 |  | 2 |  | 2
Display FOV |  | 18 |  | 30 |  | 30 |  | 30 |  | 18 |  | 25
CTDI Phantom |  | 16cm |  | 32cm |  | 32cm |  | 32cm |  | 32cm |  | 16cm
CTDI vol |  | 67.2 |  | 17.5 |  | 13.75 |  | 11.7 |  | 30.6 |  | 13.87
Notification Value |  | NA |  | NA |  | NA |  | NA |  | NA |  | NA
ACR Reference |  | 75 |  | 25 |  | -- |  | -- |  | -- |  | --
ACR Pass/Fail |  | 80 |  | 30 |  | -- |  | -- |  | -- |  | --
ACR Protocol |  | Pediatric Head |  | Pediatric Abd
RPID Code |  | RPID22 |  | RPID3
Facility Protocol
kVp
mA
Rotation Time
mAs
Effective mAs
Detector Width
# Detectors
Axial or Helical
Pitch
Table Increment
Scan FOV
Dose Reduction
Settings
Recon Kernel
Recon Width
Recon Interval
Display FOV
CTDI Phantom |  | 16cm |  | 32cm
CTDI vol
Notification Value
ACR Reference |  | 35 |  | 15
ACR Pass/Fail |  | 40 |  | 20
2. Table Travel and Return to Fixed Position Accuracy | [11 empty] | PASS
Objectives: |  | To verify that the patient table translates as indicated.
[8 empty] | Measurement |  | Tolerance |  | Result
Slice Location of the 1st Module BBs as displayed on the scanner | [7 empty] | 0 |  | -- |  | --
Slice Location of the 4th Module BBs as displayed on the scanner | [7 empty] | 120 |  | -- |  | --
Expected Table Travel Length | [7 empty] | 120 |  | -- |  | --
Slice Location of the 1st Module BBs after full extension of the table | [7 empty] | 0 |  | -- |  | --
Table Travel Accuracy | [7 empty] | 0 |  | 2 |  | Pass
Return to Fixed Position Accuracy | [7 empty] | 0 |  | 2 |  | Pass
3. Scout Prescription and Alignment Light Accuracy | [11 empty] | PASS
Objectives: |  | To verify that the scout image prescription correctly identifies the scan position.
[8 empty] | Measurement |  | Tolerance |  | Result
Slice Location of the 1st Module BBs as displayed on the scout image | [7 empty] | 0 |  | -- |  | --
Slice Location of the 1st Module BBs as displayed on the axial image | [7 empty] | 0 |  | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
Slice Location of the 4th Module BBs as displayed on the scout image | [7 empty] | 120 |  | -- |  | --
Slice Location of the 4th Module BBs as displayed on the axial image | [7 empty] | 120 |  | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
Objectives: |  | To verify that the incorporated alignment lights correctly indicate the scan position.
[8 empty] | Measurement |  | Tolerance |  | Result
Table Location of the laser alignment light of the 1st Module BBs | [7 empty] | 0 |  | -- |  | --
Slice Location of the 1st Module BBs as displayed on the axial image | [7 empty] | 0 |  | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
Table Location of the laser alignment light of the 4th Module BBs | [7 empty] | 120 |  | -- |  | --
Slice Location of the 4th Module BBs as displayed on the axial image | [7 empty] | 120 |  | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
4. Slice thickness acccuracy (ACR CT Phantom Module 1) | [11 empty] | PASS
Objectives: |  | To verify that the reconstruction slice thickness accuracy.
Slice thickness (mm) |  |  | 1 |  | 2 |  | 5 |  | 7 |  | 10
Lines detected in Top |  |  | 2 |  | 4 |  | 10 |  | 15 |  | 20
Lines detected in Bottom |  |  | 1 |  | 3 |  | 10 |  | 14 |  | 21
Measured Width (mm) |  |  | 0.75 |  | 1.75 |  | 5 |  | 7.25 |  | 10.25
5. CT Number Accuracy (ACR CT Phantom Module 1) | [11 empty] | PASS
Objectives: |  | To verify that the CT numbers reported by the CT scanner are acceptably accurate.
Material |  | Adult Head | Adult Abd | Ped Head | Ped Abd | 80 | 100 | 120 | 140 | ACR Min | ACR Max | Pass/Fail
Polyethylene |  | -- | -101.4 | -- | -- | -- | -- | -- | -- | -107 | -84 | Pass
Water |  | -4.9 | -3.8 | -- | -- | 1.48 | -3 | -3.8 | -2.5 | -7 | 7 | Pass
Acrylic |  | -- | 118.8 | -- | -- | -- | -- | -- | -- | 110 | 135 | Pass
Bone |  | -- | 881.12 | -- | -- | -- | -- | -- | -- | 850 | 970 | Pass
Air |  | -1018.95 | -994.3 | -- | -- | -994.8 | -995.31 | -994.3 | -990.6 | -1005 | -970 | Pass
6. Low Contrast Performance (ACR CT Phantom Module 2) | [11 empty] | PASS
Objectives: |  | To verify that the low-contrast performance of clinical protocols is adequate for diagnosis.
|  | Adult Head |  | Adult Abd |  | Ped Head |  | Ped Abd
Avg Disk |  | 89.76 |  | 92.5 |  | -- | -- | -- | --
Avg Background |  | 82.4 |  | 88.15 |  | -- | -- | -- | --
SD Background |  | 2.5 |  | 4.03 |  | -- | -- | -- | --
CNR |  | 2.944 |  | 1.0794 |  | -- |  | --
Tolerance |  | 1 |  | 1 |  | 0.7 |  | 0.4
7. CT Number Uniformity and Artifact Evaluation (ACR CT Phantom Module 3) | [11 empty] | PASS
Objectives: |  | To identify nonuniformities of measured CT numbers.
|  | Adult Head | Adult Abd | Ped Head | Ped Abd | 80 | 100 | 120 | 140
Top |  | -4.05 | -2.01 | -- | -- | -- | -- | -- | --
Right |  | -2.33 | -1.82 | -- | -- | -- | -- | -- | --
Botom |  | -2.42 | -1.87 | -- | -- | -- | -- | -- | --
Left |  | -3.24 | -2.4 | -- | -- | -- | -- | -- | --
Center |  | -1.46 | -0.075 | -- | -- | -- | -- | -- | --
Max Deviation |  | 2.59 | 2.325 | -- | -- | -- | -- | -- | --
Objectives: |  | To identify and correct artifacts in images of a uniform test phantom before they become severe enough to be detected in patient images.
|  | Adult Head | Adult Abd | Ped Head | Ped Abd | 80 | 100 | 120 | 140
Ring |  | No | No | -- | -- | No | No | No | No
Streak |  | No | No | -- | -- | No | No | No | No
Other |  | No | No | -- | -- | No | No | No | No
Comment |  | No | No | -- | -- | No | No | No | No
8. In-Plane Distance Accuracy (ACR CT Phantom Module 3) | [11 empty] | PASS
Objectives: |  | To verify that the system accurately measures in-plane distances of reconstructed images and adequate for diagnosis.
|  | Adult Head |  | Adult Abd |  | Ped Head |  | Ped Abd
Physical Length |  | 100 |  | 100 |  | 100 | -- | 100 | --
Measured Length |  | 100 |  | 100 |  | -- | -- | -- | --
Error% |  | 0% |  | 0% |  | -- |  | --
9. High Contrast Resolution | [11 empty] | PASS
Objectives: |  | To verify that the spatial resolution performance of clinical protocols is adequate for diagnosis.
|  | Adult Head |  | Adult Abd |  | Ped Head |  | Ped Abd |  | Hi Res Chest
Visible lp/cm |  | 7 |  | 7 |  | -- | -- | -- | -- | 8
Tolerance lp/cm |  | -- |  | 6 |  | -- | -- | -- | -- | 8
50% MTF lp/cm | [5 empty] | -- | -- | -- | --
10% MTF lp/cm | [5 empty] | -- | -- | -- | --
10. Radiation Beam Width | [11 empty] | PASS
N | T (mm) | NxT (mm) |  | Measured (mm) |  | Manufacturer Nominal (mm) |  | Manufacturer Min (mm) |  | Manufacturer Max (mm) |  | Result
2 | 0.6 | 1.2 |  | 3 |  | 3 |  | 2 |  | 4 |  | Pass
2 | 1 | 2 |  | 3.5 |  | 3.5 |  | 2.5 |  | 4.5 |  | Pass
2 | 5 | 10 |  | 12 |  | 12.5 |  | 11.5 |  | 13.5 |  | Pass
12 | 1.5 | 18 |  | 24.5 |  | 24.5 |  | 23.5 |  | 25.5 |  | Pass
16 | 1.5 | 24 |  | 30.5 |  | 30.5 |  | 29.5 |  | 31.5 |  | Pass
ACR's Recommended Performance Criteria and Corrective Action:
You may use the manufacturer’s specifications as the performance criteria. However, if these are not available, then the measured radiation beam width should be accurate to within 3 mm or 30% of the total nominal collimated beam width (N × T), whichever is greater.
11. kVp Accuracy and HVL Measurement | [11 empty] | PASS
Nominal |  | 80 |  | 100 |  | 120 |  | 140
Measured |  | 79.4 |  | 100.3 |  | 121.4 |  | 138.7
% Error |  | 0.75% |  | 0.3% |  | 1.167% |  | 0.9286%
HVL (mm) |  | 6.37 |  | 7.6 |  | 8.57 |  | 9.36
Min HVL (mm) |  | 2.9 |  | 3.6 |  | 4.3 |  | 5
12. CT Dosmetry | [11 empty] | PASS
ACR Protocol |  | Adult Head |  | Adult Abd |  | Pediatric Head |  | Pediatric Abd
RPID Code |  | RPID22 |  | RPID3 |  | RPID22 |  | RPID3
Facility Protocol |  | Routine Head |  | Abd/Pel WO
kVp |  | 120 |  | 120
mA |  | 350 |  | 500
Rotation Time |  | 1 |  | 0.5
mAs |  | 350 |  | 250
Effective mAs |  | 350 |  | 250
Detector Width |  | 1.5 |  | 1.5
# Detectors |  | 12 |  | 16
Axial or Helical |  | Axial |  | Helical
Pitch |  | 1 |  | 1
Table Increment |  | 18 |  | 24
Scan FOV |  | 30 |  | 50
CTDI Phantom |  | 16cm |  | 32cm
Displayed CTDI vol |  | 67.2 |  | 17.5
Center (mR) |  | 1323 |  | 278
Center (mR) |  | 1323 |  | 276
Center (mR) |  | 1324 |  | 278
12 O'Clock (mR) |  | 1434 |  | 587
12 O'Clock (mR) |  | 1435 |  | 588
12 O'Clock (mR) |  | 1434 |  | 588
Measured CTDIw |  | 67.5396 |  | 17.5568
Measured CTDIvol |  | 67.5396 |  | 17.5568
% Error |  | 0.5053% |  | 0.3246%
13. Soft-Copy (Monitor) Quality Control | [11 empty] | PASS
The soft copy display test evaluates the brightness, uniformity, resolution, and spatial accuracy of the display monitor(s). Using a calibrated luminance meter (Nuclear Associates Photometer LX), the luminance across the surface of the monitor is measured in units of candelas per square meter (Cd/m2) in the center and at all four corners of the display at maximum and minimum brightness levels. The maximum brightness of the monitor should exceed 90 Cd/m2 and the minimum brightness should be less than 1.2 Cd/m2. Furthermore, the values obtained at the four corners should be within 30% of the maximum brightness measured in the center of the screen. For resolution and spatial accuracy, the medical physicist will evaluate a SMPTE pattern onscreen to identify potential irregularities or performance problems.
SMPTE Pattern Evaluation
All gray level steps in the ring of gray levels visibly distinct? | [9 empty] | Pass
5% patch visible in the 0/5% patch; the 95% patch visible in the 95/100% patch? | [9 empty] | Pass
Finest line pair pattern visualized center at each of the 4 corner? | [9 empty] | Pass
Contrast Level |  | Luminance (cd/cm2)
0% |  | 0.494
10% |  | 1.72
20% |  | 4.2
30% |  | 8.9
40% |  | 15.7
50% |  | 26.3
50% |  | 26.5
60% |  | 42.2
70% |  | 63.5
80% |  | 88.6
90% |  | 122
100% |  | 168
Luminance (cd/cm2)
Center |  | 26.4
Upper Left |  | 23.5
Upper Right |  | 26.4
Bottom Right |  | 24.9
Bottom Left |  | 25
[7 empty] | Measured | Limit | P/F/NA
Maximun Luminance (cd/cm2) | [6 empty] | 168 | 100 | Pass
Minimun Luminance (cd/cm2) | [6 empty] | 0.494 | 1.2 | Pass
Luminance Ratio (LR') | [6 empty] | 340.081 | 100 | Pass
Uniformity % difference=200×(Lmax–Lmin)/(Lmax+Lmin) | [6 empty] | 11.62% | 15% | Pass