## Summary
Facility | Manufacture
Address 1 | Model
Address 2
City, State, ZIP | Gantry SN
Room Number | Clinical Engineering ID
Accreditation Information
ACR CTAP | Expiration
ACR Patient Type
Surveyor | Survey Date
Medical Physicist | Report Date
Medical Physicist Tests |  | Pass/Fail/NA
Review of CT Protocols |  | 0
Scout Prescription Accuracy |  | 0
Alignment Light Accuracy |  | 0
Table Travel Accuracy |  | 0
Radiation Beam Width |  | 0
Low-Contrast Performance |  | 0
Spatial Resolution |  | 0
CT Number Accuracy |  | 0
CT Number Uniformity |  | 0
Artifact Evaluation |  | 0
Dosimetry |  | 0
Acquisition Display Calibration |  | 0
Technologist QC Evaluation |  | Pass/Fail/NA
Water CT Number and SD (Daily)
Artifact Evaluation (Daily)
Wet Laser QC (Weekly)
Visual Checklist (Monthly)
Dry Laser QC (Monthly)
Acquisition Display QC (Monthly)
Medical Physicist's Recommendations for Quality Improvement
None.

## Data
1. CT Protocol Review
abc |  | Adult Head |  | Adult Abd |  | CT Chest WO |  | Chest w Embolus |  | C-Spine WO |  | Sinus WO
RPID Code |  | RPID22 |  | RPID3 |  | RPID16 |  | RPID147 |  | RPID21 |  | RPID62
Facility Protocol
kVp
mA |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
Rotation Time
mAs |  | 0 |  | 0 |  | 0 |  | 0 |  | 0 |  | 0
Effective mAs
Detector Width
# Detectors
Axial or Helical
Pitch
Table Increment |  | 0 |  | 0 |  | 0 |  | 0 |  | 0 |  | 0
Scan FOV
Dose Reduction
Settings
Recon Kernel
Recon Width
Recon Interval
Display FOV
CTDI Phantom |  | 16cm |  | 32cm |  | 32cm |  | 32cm |  | 32cm |  | 16cm
CTDI vol
Notification Value
ACR Reference | [5 empty] | -- |  | -- |  | -- |  | --
ACR Pass/Fail | [5 empty] | -- |  | -- |  | -- |  | --
ACR Protocol |  | Pediatric Head |  | Pediatric Abd
RPID Code |  | RPID22 |  | RPID3
Facility Protocol
kVp
mA
Rotation Time
mAs
Effective mAs
Detector Width
# Detectors
Axial or Helical
Pitch
Table Increment
Scan FOV
Dose Reduction
Settings
Recon Kernel
Recon Width
Recon Interval
Display FOV
CTDI Phantom |  | 16cm |  | 32cm
CTDI vol
Notification Value
ACR Reference
ACR Pass/Fail
2. Table Travel and Return to Fixed Position Accuracy
Objectives: |  | To verify that the patient table translates as indicated.
[8 empty] | Measurement |  | Tolerance |  | Result
Slice Location of the 1st Module BBs as displayed on the scanner | [9 empty] | -- |  | --
Slice Location of the 4th Module BBs as displayed on the scanner | [9 empty] | -- |  | --
Expected Table Travel Length | [7 empty] | 120 |  | -- |  | --
Slice Location of the 1st Module BBs after full extension of the table | [9 empty] | -- |  | --
Table Travel Accuracy | [7 empty] | 120 |  | 2 |  | Pass
Return to Fixed Position Accuracy | [7 empty] | 120 |  | 2 |  | Pass
3. Scout Prescription and Alignment Light Accuracy
Objectives: |  | To verify that the scout image prescription correctly identifies the scan position.
[8 empty] | Measurement |  | Tolerance |  | Result
Slice Location of the 1st Module BBs as displayed on the scout image | [9 empty] | -- |  | --
Slice Location of the 1st Module BBs as displayed on the axial image | [9 empty] | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
Slice Location of the 4th Module BBs as displayed on the scout image | [9 empty] | -- |  | --
Slice Location of the 4th Module BBs as displayed on the axial image | [9 empty] | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
Objectives: |  | To verify that the incorporated alignment lights correctly indicate the scan position.
[8 empty] | Measurement |  | Tolerance |  | Result
Table Location of the laser alignment light of the 1st Module BBs | [9 empty] | -- |  | --
Slice Location of the 1st Module BBs as displayed on the axial image | [9 empty] | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
Table Location of the laser alignment light of the 4th Module BBs | [9 empty] | -- |  | --
Slice Location of the 4th Module BBs as displayed on the axial image | [9 empty] | -- |  | --
Difference | [7 empty] | 0 |  | 2 |  | Pass
4. Slice thickness acccuracy (ACR CT Phantom Module 1)
Objectives: |  | To verify that the reconstruction slice thickness accuracy.
Slice thickness (mm) |  |  | 1 |  | 2 |  | 5 |  | 7 |  | 10
Lines detected in Top
Lines detected in Bottom
Measured Width (mm) |  |  | 0 |  | 0 |  | 0 |  | 0 |  | 0
5. CT Number Accuracy (ACR CT Phantom Module 1)
Objectives: |  | To verify that the CT numbers reported by the CT scanner are acceptably accurate.
Material |  | Adult Head | Adult Abd | Ped Head | Ped Abd |  |  |  |  | ACR Min | ACR Max | Pass/Fail
Polyethylene |  | -- |  | FALSE | FALSE | -- | -- | -- | -- | -107 | -84 | Pass
Water | [9 empty] | -7 | 7 | Pass
Acrylic |  | -- |  | FALSE | FALSE | -- | -- | -- | -- | 110 | 135 | Pass
Bone |  | -- |  | FALSE | FALSE | -- | -- | -- | -- | 850 | 970 | Pass
Air | [9 empty] | -1005 | -970 | Pass
6. Low Contrast Performance (ACR CT Phantom Module 2)
Objectives: |  | To verify that the low-contrast performance of clinical protocols is adequate for diagnosis.
|  | Adult Head |  | Adult Abd |  | Ped Head |  | Ped Abd
Avg Disk
Avg Background
SD Background
CNR |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
Tolerance |  | 1 |  | 1 |  | 0.7 |  | 0.4
7. CT Number Uniformity and Artifact Evaluation (ACR CT Phantom Module 3)
Objectives: |  | To identify nonuniformities of measured CT numbers.
|  | Adult Head | Adult Abd | Ped Head | Ped Abd | 0 | 0 | 0 | 0
Top |  |  |  | FALSE | FALSE | -- | -- | -- | --
Right |  |  |  | FALSE | FALSE | -- | -- | -- | --
Botom |  |  |  | FALSE | FALSE | -- | -- | -- | --
Left |  |  |  | FALSE | FALSE | -- | -- | -- | --
Center |  |  |  | FALSE | FALSE | -- | -- | -- | --
Max Deviation |  | 0 | 0 | 0 | 0 | -- | -- | -- | --
Objectives: |  | To identify and correct artifacts in images of a uniform test phantom before they become severe enough to be detected in patient images.
|  | Adult Head | Adult Abd | Ped Head | Ped Abd | 80 | 100 | 120 | 140
Ring
Streak
Other
Comment
8. In-Plane Distance Accuracy (ACR CT Phantom Module 3)
Objectives: |  | To verify that the system accurately measures in-plane distances of reconstructed images and adequate for diagnosis.
|  | Adult Head |  | Adult Abd |  | Ped Head |  | Ped Abd
Physical Length
Measured Length
Error% |  | 0% |  | 0% |  | 0% |  | 0%
9. High Contrast Resolution
Objectives: |  | To verify that the spatial resolution performance of clinical protocols is adequate for diagnosis.
|  | Adult Head |  | Adult Abd |  | Ped Head |  | Ped Abd |  | Hi Res Chest
Visible lp/cm
Tolerance lp/cm |  | -- |  | 6 |  | FALSE |  | FALSE |  | 8
50% MTF lp/cm | [5 empty] | FALSE |  | FALSE
10% MTF lp/cm | [5 empty] | FALSE |  | FALSE
10. Radiation Beam Width
N | T (mm) | NxT (mm) |  | Measured (mm) |  | Manufacturer Nominal (mm) |  | Manufacturer Min (mm) |  | Manufacturer Max (mm) |  | Result
2 |  | 0 | [9 empty] | Pass
2 |  | 0 | [9 empty] | Pass
2 |  | 0 | [9 empty] | Pass
12 |  | 0 | [9 empty] | Pass
16 |  | 0 | [9 empty] | Pass
ACR's Recommended Performance Criteria and Corrective Action:
You may use the manufacturer’s specifications as the performance criteria. However, if these are not available, then the measured radiation beam width should be accurate to within 3 mm or 30% of the total nominal collimated beam width (N × T), whichever is greater.
11. kVp Accuracy and HVL Measurement
Nominal |  | 0 |  | 0 |  | 0 |  | 0
Measured
% Error |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
HVL (mm)
Min HVL (mm) |  | 2.9 |  | 3.6 |  | 4.3 |  | 5
12. CT Dosmetry
ACR Protocol |  | Adult Head |  | Adult Abd |  | Pediatric Head |  | Pediatric Abd
RPID Code |  | RPID22 |  | RPID3 |  | RPID22 |  | RPID3
Facility Protocol |  | Routine Head |  | Abd/Pel WO
kVp |  | 120 |  | 120
mA |  | 350 |  | 500
Rotation Time |  | 1 |  | 0.5
mAs |  | 350 |  | 250
Effective mAs |  | 350 |  | 250
Detector Width |  | 1.5 |  | 1.5
# Detectors |  | 12 |  | 16
Axial or Helical |  | Axial |  | Helical
Pitch |  | 1 |  | 1
Table Increment |  | 18 |  | 24
Scan FOV |  | 30 |  | 50
CTDI Phantom |  | 16cm |  | 32cm
Displayed CTDI vol |  | 67.2 |  | 17.5
Center (mR)
Center (mR)
Center (mR)
12 O'Clock (mR)
12 O'Clock (mR)
12 O'Clock (mR)
Measured CTDIw |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
Measured CTDIvol |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
% Error |  | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
13. Soft-Copy (Monitor) Quality Control
The soft copy display test evaluates the brightness, uniformity, resolution, and spatial accuracy of the display monitor(s). Using a calibrated luminance meter (Nuclear Associates Photometer LX), the luminance across the surface of the monitor is measured in units of candelas per square meter (Cd/m2) in the center and at all four corners of the display at maximum and minimum brightness levels. The maximum brightness of the monitor should exceed 90 Cd/m2 and the minimum brightness should be less than 1.2 Cd/m2. Furthermore, the values obtained at the four corners should be within 30% of the maximum brightness measured in the center of the screen. For resolution and spatial accuracy, the medical physicist will evaluate a SMPTE pattern onscreen to identify potential irregularities or performance problems.
SMPTE Pattern Evaluation
All gray level steps in the ring of gray levels visibly distinct?
5% patch visible in the 0/5% patch; the 95% patch visible in the 95/100% patch?
Finest line pair pattern visualized center at each of the 4 corner?
Contrast Level |  | Luminance (cd/cm2)
0%
10%
20%
30%
40%
50%
50%
60%
70%
80%
90%
100%
Luminance (cd/cm2)
Center |  | #DIV/0!
Upper Left
Upper Right
Bottom Right
Bottom Left
[7 empty] | Measured | Limit | P/F/NA
Maximun Luminance (cd/cm2) | [6 empty] | 0 | 100
Minimun Luminance (cd/cm2) | [6 empty] | 0 | 1.2
Luminance Ratio (LR') | [7 empty] | 100
Uniformity % difference=200×(Lmax–Lmin)/(Lmax+Lmin) | [6 empty] | #DIV/0! | 15%